## Slide 1: Capabilities of Hybrid SIMS-SSAMS System for Nuclear Forensics Applications
K. S. Grabowski, K. C. Fazel, D.L. Knies*
US Naval Research Laboratory
Washington, DC 20375
*Now at Coolescence, Boulder CO
Notes: ETH U measurement use He stripper
“At ion energies of 330 keV a maximum transmission of about 40% is observed for
238U and 232Th [51] when using He as stripper gas. In the case of
236U measurements a higher stripper thickness has to be applied
to achieve a sufficiently high molecular dissociation rate [50]
resulting in a transmission of about 35% under real measurement
conditions. For Pu and Pa measurements the interfering UH and
ThH molecules can be efficiently suppressed by chemical preparation
so that the stripper gas thickness can be adjusted for optimal
transmission.”
References:
Leibniz Institute for Solid State and Materials Research Dresden http://www.ifw-dresden.de/institutes/ikm/organisation/dep-31/methods/secondary-ion-mass-spectrometry-sims
Benninghoven, A., Rudenauer, F. G., & Werner, H. W. (1987). Secondary Ion Mass Spectrometry: Basic Concepts, Instrumental Aspects, Applications and Trends (1st ed.). Wiley-Interscience.
Jacob, S. A. W., Suter, M., & Synal, H.-A. (2000). Ion beam interaction with stripper gas – Key for AMS at sub MeV. Nuclear Instruments and Methods in Physics Research Section B: Beam Interactions with Materials and Atoms, 172(1–4), 235–241. doi:10.1016/S0168-583X(00)00205-6
Suter, M. (1998). A new generation of small facilities for accelerator mass spectrometry. Nuclear Instruments and Methods in Physics Research Section B: Beam Interactions with Materials and Atoms, 139(1–4), 150–157. doi:10.1016/S0168-583X(97)00945-2 (shows method of extrapolating high to low E)
Christl, M., Wacker, L., Lippold, J., Synal, H.-A., & Suter, M. (2007). Protactinium-231: A new radionuclide for AMS. Nuclear Instruments and Methods in Physics Research Section B: Beam Interactions with Materials and Atoms, 262(2), 379–384. doi:10.1016/j.nimb.2007.05.017
Christl, M., Vockenhuber, C., Kubik, P. W., Wacker, L., Lachner, J., Alfimov, V., & Synal, H.-A. (n.d.). The ETH Zurich AMS facilities: Performance parameters and reference materials. Nuclear Instruments and Methods in Physics Research Section B: Beam Interactions with Materials and Atoms, (0). doi:10.1016/j.nimb.2012.03.004
BETZ, H.-D. (1972). Charge States and Charge-Changing Cross Sections of Fast Heavy Ions Penetrating Through Gaseous and Solid Media. Reviews of Modern Physics, 44(3), 465–539. doi:10.1103/RevModPhys.44.465

## Slide 2: atomic ion SIMS (ai-SIMS) at NRL

## Slide 3: Instrument Overview – A World Unique Tool
Achieve molecular-interference-free trace analysis with spatial resolution
Combine strong points of Secondary Ion Mass Spectrometry (SIMS) and Accelerator Mass Spectrometry (AMS) into one instrument
Uses:
Nuclear forensics
      Environmental samples of particulates
      Fingerprint of source material
Advanced Materials Science needs
        Structural
        Electronic
        Optical
        Biological
        NASA Discovery Mission samples

## Slide 4: SIMS Proven tool Spatial information Limited by molecular interferences (M>90)
Mass (atomic mass units)
Intensity (counts)
Resolution Needed (M/DM)
Mass (atomic mass units)
1 10 102 103 104 105 106 107
90 100 110 120 130 140 150

## Slide 5: AMS Ultra-trace capable (14C/12C > 10-15) Molecules destroyed Limited spatial information Negative ions only
104 better with + ions
Actinides
Lanthanides
Sensitivity

## Slide 6: A Cornerstone Technique
+ ions
- ions
Molecular
Ions
No
Molecular
Ions
~µm Resolution
~mm Resolution
ICPMS
AMS
+ SIMS-SSAMS
- SIMS-SSAMS
+ SIMS
- SIMS

## Slide 7: Recent Developments ® ai-SIMS Now
AMS
  NRL Trace Element AMS (TEAMS) results
    Negative ion injection yields poor sensitivity
    Detection of +1 ions preferred, to remove ambiguities
  New smaller Single Stage AMS available and proven in radiocarbon analysis
    Enables injection of either + or - ions
    Abundant +1 ions for detection
Opportunity
  Employ SIMS to generate preferred + or - ions
  Inject + or - ions into small bipolar accelerator, detect +1 ions

## Slide 8: Implementation of ± ai-SIMS at the Naval Research Laboratory (NRL) in Washington DC

## Slide 9: Flow Diagram of ai-SIMS
Ion Source
Sample
ESA
Magnet
Beam Switch
Injection Magnet
Analyzing Magnet
ESA
Detection
Gas Cell
300 kV Accelerator
CAMECA 4f
NEC SSAMS

## Slide 10: NEC SSAMS Drawing
28.3’
±300 kV Deck
NRL injector and CAMECA 4f here
Molecule Destruction Chamber
Detectors
Magnet
Electrostatic Analyzer (ESA)
(2) Motor/Generator Sets for AC Power on Deck
Injection Magnet
Acceleration Tube

## Slide 11
Molecule Destruction Chamber
Magnet
Electrostatic Analyzer (ESA)
Acceleration Tube
Insulation to “Bounce” Energy

## Slide 12: Interchangeable “Stack” for ±300 kV

## Slide 13: Cameca ims 4f, Coupling to NEC SSAMS, and Control Station

## Slide 14: Instrument Development

## Slide 15: Charge State
Gas chamber
+1,2,3,…
±1

## Slide 16: Want Charge State +1 Ions for Analysis
 | Injected | 
Species | 118Sn2 | 236U
Charge State (q) | +1 | +1
E (keV) | 300 | 300
M (u) | 236 | 236
 | Analyzed | 
Species | 118Sn | 236U
Charge State (q) | +1 | +2
E (keV) | 150 | 300
E/q | 150 | 150
M (u) | 118 | 236
M/q | 118 | 118
To avoid molecule breakup interferences

## Slide 17: Equilibrium Conditions for Low-energy Heavy-ion Charge States Needed
Betz,H-D., Charge states and charge-changing cross sections of fast heavy ions penetrating through gaseous and solid media. Rev. Mod. Phys. 44 (1972) 465. Copyright (1972) by the American Physical Society.
Needed for each ion, energy, and gas used
Model based on charge changing cross sections can be developed
Injected Distribution
Equilibrium Distribution

## Slide 18: Choose Gas and Energy to Maximize U1+
Vockenhuber, C. et al. The potential of He stripping in heavy ion AMS. Nucl. Instrum. Meth. B 294 (2013) 382. Reprinted with permission of Elsevier.
Ar, Kr, Xe, H2 better than He?
Desired

## Slide 19: Molecule Dissociation and Beam Broadening
Gas chamber
1+

## Slide 20: 63Cu2
Ar breakup gas reduced 63Cu2 intensity by 107, while reducing 63Cu intensity by 70%. Residual 126Te may appear from under 63Cu2 molecule ion peak.

## Slide 21: Cross Sections
Molecular ion destruction cross section ~20x that of atomic ion attenuation. Functional dependence on energy, gas, and molecule will be used to develop a predictive model to optimize methodology. Sparse data available at these conditions.

## Slide 22: Mass Distribution of Hf Isotopes from HfH2 Sample
Hf-isotope mass distribution perturbed by Hf hydride molecules.
Use of 0.45 sccm flow of Ar improves measurement to near natural abundance values.

## Slide 23: Nuclear Forensics Concerns
Monohydrides can obscure measurement of pertinent isotopes:
Neutron exposure of U
  236U, 239Pu
Age of processed U
  231Pa, 210Pb
Bomb component
  210Po, 239Pu/240Pu
Repurposing of 232Th
  233Pa/233U

## Slide 24: Instrument Applications

## Slide 25
Environmental samples of particulates
  Need isotope ratios: 236/238, 235/238
  Hydride molecules formed (235UH ~ 236U)
  Abundance sensitivity issue with 238U
  SIMS currently used, precision limitations
Fingerprint of source material
  236U present
  Lanthanide levels present in natural U
  Pu isotopes in mixed U/Pu particles
Nuclear Forensics and Safeguard Verification

## Slide 26: Plan
Complete control integration and operational testing
Model and test performance characteristics of ai-SIMS as function of ion charge state, energy, stripper gas and pressure for various materials
Initiate measurement of representative samples
Determine sensitivity and precision for prototype samples
Measure isotope ratio and elemental abundance of standards, then real samples

## Slide 27: Summary
Bipolar ai-SIMS =
  First in world
  104 improved sensitivity for Lanthanides, Actinides
  Removes molecular interferences for M > 90 u
  Provides image and depth profile composition
  Challenges exist, but recent high-mass research shows path forward
  Has many relevant applications
  Installed at NRL, undergoing “shakedown”
Welcoming samples for analysis and collaborations