## Slide 1
Enhancement of Peak to Compton Ratio (P/C) using a New Array
Design for Safeguards Applications
R.A. El-Tayebany1, A.A.Hamed1, A.Abdelsalam2,W. Osman2, Sayed El-Mongy1
1 Egyptian Nuclear and Radiological Regulatory Authority(ENRRA), Cairo, Egypt
2 Physics Department, Faculty of Science, Cairo University, Cairo, Egypt
reltyebany@ymail.com
This paper focuses on installation of a new geometry and array of detectors to investigate reduction and suppression of unwanted noises and background to enhance P/C ratio.
The new array consists of three sodium iodide (NaI) detectors; one of them is in annular perpendicular position and the others are guards surrounding the main Hyper Pure Germanium Detector (HpGe).
The optimum configuration for the array was selected to maximize P/C ratio and minimize the noises and Compton continuum that produced at higher-energy gamma-rays.
The enhancement of the Peak-Compton (P/C) ratios was investigated using the radioactive source 137Cs. It was observed that the new array configuration enhanced the P/C ratios compared to single HPGe detector in range 387.6±6.12 to 1001.12±7.4.
The design was also investigated and applied on standard nuclear materials (SNM) at the photo peak 185.7 KeV of 235U. The system gave great value in the net area for photo peak 185.7 KeV for investigated SNMs samples. Since, The maximum difference was found to be 17.6%. The results are given, discussed and interpreted.
ABSTRACT
It is necessary to understand the physical explanation for the detector array set up that contains HpGe detector which is surrounded by two guard detectors and one annular, usually NaI. The HpGe and NaI detectors are operated in anti-coincidence mode. The photons that are escaped from HpGe detector are interacted and deposited within the NaI detectors to form Compton continuum that can be subtracted from the spectrum to reduce the un interested background [6].
To investigate and validate the new array of detectors, the SNMs are used with different enrichment. The net area of photo peak at energy line 185.7 KeV is measured at different distances of the sample from the HpGe detector The net area of energy line 185.7KeV of 235U is registered in live time 900 S. the energy range of photo peak 185.7 KeV is taken from 178 KeV to 197 KeV. The results are compared with the result obtained using single HpGe-detector in the experimental results.
OUTCOME
In gamma-ray spectrometry measurements, the predominant mode of photon interaction depends on the energy of the incident photons and the atomic number of the target material.
Gamma spectrometry with high-purity germanium (HPGe) detectors is routinely used for the Non- Destructive assay techniques (NDA). Improvements to the NDA can be achieved by minimizing the Compton continuum, through the use of array of detectors.
The non-destructive identification of nuclear material samples can be achieved using different gamma spectrometry techniques for accurate verification of NMs. However, the sensitivity of Gamma spectroscopy is dependent on many factors. This includes the resolution and efficiency of the detector, and the amount of background radiation seen.
BACKGROUND
Sources and detectors assembly Design
The experiments were carried out at different distance between the HpGe-detector and the sample at 1,2,3 and 4 cm since, the guard detectors kept with fixed position close to the active volume of HpGe detector in each case. In The first experiment, the radioactive source 137Cs with activity of 5μCi (Certified Standard point sources [Eckert & Ziegler, California-USA (2006); with total uncertainty ±3.0%]) was used in the presence of single HpGe detector alone. The photo peak was collected in the energy range 650-670 KeV marked by ROI (Region of Interest) in MAESTRO software. The net area of photo peak at the energy line 661.7 KeV was registered in live time of 900 sec. in each distance. The same experiment was repeated with marked another area (Compton continuum) which is in the energy ranges 430-460 KeV to estimate P\C ratio. These steps were repeated at each source-to-detector distance to calculate the P\C ratios.
The above procedures were executed for certain array where the HpGe detector surrounded by two NaI guard detectors and one annual,as shown in figure (1),which are operated with WINSPEC software in the same time of HpGe detector to determine P\C ratio outside and inside the shield and compare the results. The detectors specifications are displayed in table (1).
Table (1).The used Detector Specifications
CHALLENGES / METHODS / IMPLEMENTATION
Title of the graphs/figures
This work aims at the investigation and improvement of detection accuracy for several safeguards key elements using new detectors assembly. The experimental work showed that the P/C values were enhanced using the array detectors design with respect to single HPGe usage only from 387.6±6.12 to 1001.12±7.4. This result indicates and proves that the proposed new detectors array design achieved the target. The system is quite capable for the assay of small amounts and depleted nuclear materials and radioactive source with relatively high accuracy and enhanced P/C ratios. The system gives great value in the net area for photo peak 185.7 KeV for investigated SNMs samples.Since, The difference ranges from 0.91% to 17.55%. The designed system is being developed to be comparable with the commercial Compton suppression systems.
CONCLUSION
ID: 82
This research has been carried out in Nuclear Safeguards Laboratory- Nuclear Safeguards& physical protection department of EgyptianNuclear & Radiological Regulatory Authority (ENRRA).
ACKNOWLEDGEMENTS / REFERENCES
Chart: 
| Single Ge detector | non - shielded Array | shielded Array
Distance = 1cm | 395.4495 | 784.8456 | 962.9783
Distance = 2cm | 364.2675 | 939.6862 | 929.8067
Distance = 3cm | 387.5766 | 928.6018 | 950.9592
Distance = 4cm | 439.4646 | 839.9284 | 1001.1185
Fig.(2): The experimental results for P\C ratio with different distance between the HpGe detector and the radioactive point source 137Cs in all the conditions.
Fig.(4): Net area at energy line 185.7 KeV versus the distance between the HpGe detector and the SNM (Depleted Uranium) in all conditions.