## Slide 1
Numerical Method for High Count-Rate Dead-Time Correction in Neutron Multiplicity Counting using Multi-Channel List-Mode Recorders
L. HOLZLEITNER
European Commission – Joint Research Centre (JRC), Karlsruhe, Germany
E-mail: Ludwig.Holzleitner@ec.europa.eu
D. HENZLOVA, M.T. SWINHOE, V. HENZL
NEN-1 Safeguards Science Technology Group, Los Alamos National Laboratory, Los Alamos, USA

## Slide 2
MC list-mode Counter
Multiplicity
Evaluation
3He Tube
Preamplifier
Polyethylene
Sample cavity
Office
Data-Link
Traditional neutron counters/detectors sum up the signals from different preamplifiers, possibly using a de-randomizer.
Multi-channel list mode recorders provide both accurate time and channel information of recorded signals. Recording signals from individual preamplifiers using such devices provide additional possibilities for data analysis and other advantages. One of it is advanced dead-time correction as described here.
A neutron detector in safeguards has a body of polyethylene with a cavity for the sample. Around it, gas proportional counters, often 3He tubes, are embedded in the polyethylene body.
De-Randomizer replaced by
Multi-Channel List-Mode recorder
Detektor

## Slide 3
(eqn.7)
Dead-time loss can be estimated by repeatedly solving the matrix - eqn. 7 for evert time and for every meaningful (means non-zero) right hand  second, estimated pulse train.
Lost pulses on channels are estimated from observed ones on all other channels. This is done simultaneously (using matrix equat.) since probabilities of pulse arrival are interlinked!
… global time of the detector
… time at certain channels from a specific event (leading pulse)
count is a observed pulse accounted for at channel i and time t.
relative efficiency of channel i where of course .
probability (betw. 0 - 1) for losing a pulse at channel has as time , time from the last recorded pulse on channel i.

## Slide 4
Multiplicity histograms can be corrected by sorting-in lost pulses using statistical methods.
This results in corrected Singles / Doubles / Triples.
(a) Original count rate 0.52 M c/s
(d) Original count rate 4.19 M c/s
(c) Original count rate 2.09 M c/s
(b) Original count rate 1.05 M c/s
MCNP simulations show that the estimation of lost pulses is quite precise;
The results for multiplicity correction by sorting-in of lost pulses are reasonable, however it still needs some refinement.
Dead-time probabilities can be estimated using a iterative process involving eqn. 7.
Notes:   New approaches to all individual links of the traditional neutron counting chain from the detector, to signal processing, to a interpretation.
  The objective is to report results to ENER along the way, and demonstrate integration of all new developments in a high efficiency counter in PERLA by 2020.

## Slide 5
Rossi-Alpha distribution first (4.5 µs) from MCNP simulation
Advantage over other methods:
No prior calibration necessary: is self-calibrating from measurement data.
Suitable for very high count-rates
Transfer of channel-information provides increased diagnostics capabilities
Compensates for double pulsing is as long as it does not exceed dead-time loss. Double-pulsing correction could be built in future using a similar technique.
Original count rate 1.05 M c/s:
Using multi-channel list-mode recorders:

## Slide 6
Questions?
Thank you!
Ludwig.Holzleitner@ec.europa.eu