## Slide 1: International standards for the performance of radiation detection instruments used in the global nuclear security framework International Conference on Nuclear Security ICONS2020, 10-14 February 2020, IAEA (Vienna)
Miroslav VOYTCHEV, IRSN (France), Secretary of IEC/SC45B
Radoslav RADEV, LLNL (US), Chairman of IEC/SC 45B
Radiation protection instrumentation
LLNL-PRES-799397
This work performed under the auspices of the U.S. Department of Energy by Lawrence Livermore National Laboratory under Contract DE-AC52-07NA27344

## Slide 2: IEC - International Electrotechnical Commission
Founded in 1906
Prepares and publishes international standards for all electrical, electronic and related technologies
National Committees (NC)
    62 member countries
    23 associated member countries
Technical Committees (TC) and Sub-committees (SC)
    207 TCs and SCs
    more than 560 WGs

## Slide 3: IEC/SC 45B Radiation protection instrumentation
Prepare standards that address instrumentation used for:
  the measurement of ionizing radiation in the workplace, to the public and in the environment for radiation protection purposes
  illicit trafficking detection and identification of radionuclides
  radiation-based security screening
Country membership
  Participating countries (P-members): 20, Observer countries (O-members): 15
More than 100 experts
  Testing laboratories
  Governmental agencies
  Manufacturers
  Users
Standards and technical documents
  Published: 57, In development: 10
Liaisons with ISO, IAEA, ICRP, ICRU…
7 Working Groups (WGs)
  WG 15: Illicit Trafficking Control Instrumentation using spectrometry, personal electronic dosimeter and portable dose rate instrumentation

## Slide 4: Standards on instrumentation used for detection of illicit trafficking of radioactive material
11 international standards published since 2003
Revisions undertaken since 2014
  Harmonization
  Taking into account the feedback from different testing programs and international projects as ITRAP+10
Types of instruments
  Body-worn
  Portable or hand-held
  Portal
  Vehicle-mounted
Data output format standard
  IEC 62755 (2012)

## Slide 5: Body-worn
IEC 62401 Ed. 2 (2017) Alarming Personal Radiation Devices (PRD)
IEC 62618 Ed. 2 (in progress) Spectroscopy-Based Alarming Personal Radiation Devices (SPRD)
IEC 62694 Ed.2 (in progress) Backpack Based Radiation Detector (BRD)

## Slide 6: Portable or hand-held
IEC 62327 Ed. 2 (2017) Hand-held Radionuclide Identification Devices (RID)
IEC 62533 (2010) Hand-held highly sensitive Photon Devices (GSD)
IEC 62534 (2010) Hand-held highly sensitive Neutron Devices (NSD)

## Slide 7: Portal
IEC 62244 Ed. 2 (2019) Portal Monitors (RPM)
IEC 62484 Ed. 2 (in progress) Spectroscopy-Based Portal Monitors (SPRM)

## Slide 8: Vehicle-mounted
IEC 63121 (in progress) Vehicle-mounted mobile systems (VMS)

## Slide 9: Basic radiation detection requirements
 | Max rate of false alarm | Gamma or neutron alarm | Gamma dose rate indication error | Gamma dose rate over-range | Additional tests
PRD | 1 alarm in 1 h | ≤ 2 s at increase of 0.5 µSv∙h–1 above background | ≤ ±50 % | At 2 times max range | 
SPRD | 3 alarms in 8 h | ≤ 3 s at increase of 0,5 µSv∙h–1 above background | ≤ ±50 % | At 2-10 times max range or 1 mSv∙h–1 | 
BRD | 1 alarm in 1 h | ≤ 2 s with gamma fluence rate of 4 s-1·cm | ≤ ±30 % | At 1.5 times max range | Angular dependence
RID | - | ≤ 1 s at increase of 0.5 µSv∙h–1 above background | ≤ ±30 % | At 1.5 times max range | Gamma source localization
GSD | 1 alarm in 1 h | ≤ 3 s at increase of 0,05 µSv∙h–1 above background | ≤ ±30 % | At 10 times max range or 1 mSv∙h–1 | 
NSD | 1 alarm in 1 h | ≤ 2 s at exposure from 252Cf to 0.1 n∙s–1.cm–2 | - | At 10 times max range | 
RPM | 2 alarms in 10 h | At radiation level greater than a threshold | - | At 1.5 times max range | Warning when backgr. change
SRPM | 1 alarm in 10 h | Alarm at radionuclide identification | - | At 1.5 times max range | Warning when backgr. Change
VMS | 1 alarm in 2 h | At radiation level greater than a threshold | - | At 1.5 times max range | Background effects

## Slide 10: Other radiation tests
Detection of gradually increasing radiation
Neutron detection alarm (if applicable)
Personal protection alarm
Radionuclide identification (RID)
  Single: 241Am, 133Ba, 60Co, 137Cs, 67Ga, 131I, 99mTc, 40K (KCl or KOH), 201Tl, 226Ra, 232Th, 238U (DU), 235U (HEU) and 239Pu (WGPu)
  Mixed: 131I + WGPu, NORM+HEU, NORM+WGPu

## Slide 11: Climatic, mechanical and EMC requirements and tests
  Climatic: ambient temperature, temperature shock, low/high temperature start-up, relative humidity, moisture and dust protection
  Mechanical: drop, vibrations, microphonics/impact and mechanical shock
  Electromagnetic: electrostatic discharge, radio frequency (RF) immunity, radiated emissions, magnetic fields and AC line powered equipment requirements
  Electrical: battery lifetime and power requirements
  Documentation

## Slide 12: Conclusion
Criteria and compliance test methods in the standards
    Optimization
    Compromise
    Consensus
Objects of the standards
    Describe the performance requirements and functional criteria
    Develop testing methods for evaluating the performance of the applicable instrumentation
    Specify general characteristics, general test procedures, radiation characteristics, climatic, mechanical, electromagnetic and electric characteristics
Compliance with standard requirements provides
    Manufacturers with internationally acceptable specifications
    Device users with assurance of the rigorous quality and accuracy of the measurements
International use
    CENELEC transposition as EN standards
    Harmonization with US ANSI standards by the Accredited Standards Committee “Radiation Detection”

## Slide 13: Thank you for your attention. Any questions?
For contacts:
miroslav.voytchev@irsn.fr
radev1@llnl.gov