- Background and Goal of the present work

Selective internal radiation therapy (SIRT) using Yttrium-90 (90Y) labelled microsphere is increasingly used to treat hepatocellular carcinoma (HCC) and metastatic hepatic tumour. Nowadays, two types of microspheres are available for clinical practices: 90Y labelled glass-microspheres (TheraSphere; MDS Nordion, Ottawa, Canada) and 90Y labelled resin microspheres (SIR-sphere; SIRTEX, Sydney, Australia)[1-3]. For 90Y SIRT procedure, the treatment is required to process by an intervention radiologist, radiation technologist and nurse. 90Y microsphere manufacturer recommends to indirectly handle radiotracer by using the radiation shielding apparatus produced from polymethyl methacrylate (PMMA) or also known as acrylic. However, many studies were reported that bremsstrahlung radiation can still be produced from the low atomic number materials while performing 90Y SIRT procedure [4, 5-8]. This might lead to increase of radiation exposure to radiation workers for 90Y SIRT. The primary aim of this work was to determine the extent of the 90Y bremsstrahlung radiation produced from the PMMA radiation shielding apparatus. In addition, the radiation safety aspect of 90Y SIRT was investigated by Monte Carlo simulation.

- Materials and Methods

2.1. Monte Carlo Simulation with MCNP

MCNP5, Monte Carlo N-Particle 5, is a three-dimension computational transport code that can be used practically for all particles and all energies [9]. Applications for the code are quite board, including detector design, radiation dosimetry and radiation safety. In this work, MCNP5 is relied on to simulate 90Y bremsstrahlung radiation produced from shielding apparatus through the user-defined geometry and material compositions. MCNP is set up in photon and electron mode with default physics parameters to enable MCNP to track down all source electrons and photons produced from subsequent bremsstrahlung radiation. The discrete energy spectrum of electrons source defined in the MCNP simulation was based on beta decay of 90Y that compose of 92.4 keV (0.0000014%), 518.0 keV (0.017%) and 2278.7 keV (99.983%)[10]. Both photons and electrons were simulated and followed when their energies were above 1 keV. The ENDF/B-IV library is used for all atomic and cross section data [11]. Several tally options are available in the MCNP5. The F8 tally was used in this study because it provides the bremsstrahlung energy distribution with pulse unit [12]. Relative errors of MCNP5 results in this work were limited within 10% by sufficiently increasing numbers of simulating particles.

  - Validation of MCNP with Experimental Measurement

MCNP is so versatile that it provides a user with an ability to capture effects from energy broadening and to describe the full width half maximum (FWHM) of the detector through specifying several parameters in the Gaussian Energy Broadening (GEB) and special treatment for tallies FT input cards, respectively. The GEB option was used to simulate a physical radiation detector which is broadening tallies energy into Gaussian distribution.

(1)

Where E, E0, C and A are the broadened energy, unbroadened energy of the tally, normalization constant and Gaussian width. The latter parameter is related to the full width half maximum (FWHM) that the desired FWHM is specified by the user–provided constants, a, b, and c as illustrated in equation 2 [13]

(2)

Where E is the photon energy in MeV. In this study, three gamma sources (241Am, 137Cs and 60Co) were used to obtain the FWHM data for determining a, b and c parameters [13]. For our work, the constants a, b and c used in equations were -9.978E-3 MeV, 8.9944E-2 MeV1/2 and -3.87858E-1 MeV-1 respectively.

In order to obtain appropriate parameters for MCNP to subsequently provide reliable assessments on the equivalent dose rates on the medical staff, MCNP must be validated against experimental measurements.  An experimental setup had a 2-inch × 2-inch thallium-doped sodium iodide (NaI(Tl)) scintillation detector (Mirion Technologies Inc., California, USA) placed in front of a 137Cs standard point source (figure 1).

Figure 1. Experimental set-up of scintillation   Figure 2. MCNP5 geometry of 137Cs measured with

detector with 137Cs standard source               NaI(Tl) detector for code simulation

The detector was previously calibrated before performing all measurements. The MCNP5 input for the NaI(Tl) scintillation detector including crystal, housing and photomultiplier tube (obtained from the manufacturer specification), 137Cs point source and the related geometry were created to resemble the experiment (figure 2).

2.3 Simulation of 90Y Bremsstrahlung Photons Spectrum and Safety Aspect

The PMMA shielding apparatus used in SIRT was simulated together with PMMA vial similar as clinical setting. 90Y bremsstrahlung photons spectrum produced from radiation shielding apparatus with NaI(Tl) detector was simulated (figure 3).

Figure 3  MCNP5 geometry for of 90Y with the

radiation shielding apparatus with NaI(Tl) detector

Figure 4  Simplified geometry model of clinical SIRT condition using data given in ICRP publication 89 (figure 4a), the reference phantom positioned at 30 cm from radiation shielding apparatus (figure 4b), and the finger model on the knob (figure 4c)

To estimate the safety aspects of radiation workers during SIRT, the activity of 3 GBq was simulated. This selected activity was based on the maximum activity according to the empiric planning of 90Y resin microsphere [14]. The simulation was based on the 90Y source placed in the radiation shielding apparatus and then positioned at a 30 cm distance from the simplified geometry phantom model (from the information given in ICRP (International Commission on Radiation Protection) publication 89 on the basic anatomical and physiological data for use in radiological protection: reference values) [15]. The simulated reference phantom was homogeneous tissue equivalent material (MS20) with density of 1.00 g/cm3. The finger equivalent dose was simulated when the finger model (diameter of 1 cm. and length of 5 cm.) contacted to the knob of the apparatus (figure 4). The equivalent dose rate (EDR) to the head, neck, body, and lower extremity from bremsstrahlung radiation produced from the PMMA radiation shielding apparatus with the simplified geometry model were simulated in the Monte Carlo simulations. The EDR is calculated by multiplying the tally particle fluence with the photons flux to dose rate conversion factor from ICRP publication 21 and the source activity [16].

- Results and Discussions

NaI(Tl) scintillation detector with 137Cs standard source was used to validate the MCNP5 code. The GEB option was applied to better simulate a physical radiation detector in which energy peak exhibit Gaussian energy broadening. Figure 5 shows the count rates obtained from the experimental NaI(Tl) detector and the MCNP simulation with GEB. The total count rates and simulated count rates were aligned in good agreement with a difference of 0.14%. 137Cs standard source has been chosen to verify Monte Carlo simulations compared with the radiation detector measurement in a number of studies due to it is widely used to calibrate NaI(Tl) scintillation detector [17,18].

Figure 5. 137Cs count rates obtained from the      Figure 6. Simulated spectrum of the 90Y bremsstrahlung

experimental measurement with NaI(Tl)                radiation with PMMA radiation shielding apparatus

detector and MCNP simulation with GEB

Figure 6 illustrates the simulated spectrum of the 90Y bremsstrahlung radiation with PMMA radiation shielding apparatus. The spectrum showed a continuous and broad photon spectrum with peak at 40 keV. Our finding was in line with previous result which performed Monte Carlo simulation (MCNP6) of 90Sr/90Y source with Plexiglas (another trade name of PMMA) [19]. The reported peak from previous study was at 38 keV in the case of Plexiglas. This study also indicated that the production of bremsstrahlung radiation from Plexiglas shielding of 90Sr/90Y was much greater than bremsstrahlung radiation that produced by 32P and 89Sr. Hence, bremsstrahlung radiation dose resulting from the usage of 90Y for SIRT should be considered. Also, there is room for improvement on the materials for 90Y shielding for example barium borate glasses shielding [20, 21].

To obtain further information on radiation safety aspect, the EDR to head, neck, body, and lower extremities to the radiation worker using simplified geometry model using data given in ICRP publication 89 was simulated with 90Y activity of 3 GBq, at distance of 30 cm. Consequently, the finger dose was simulated based on the assumption that the intervention radiologist finger contracted to the knob of the radiation shielding apparatus. The EDR are tabulated in table 1.

| Organ | EDR (µSvh-1) |
| --- | --- |
| Head | 4.9 ± 0.6 |
| Neck | 6.2 ± 0.1 |
| Body | 18.9 ± 0.4 |
| Lower Extremity | 13.1 ± 0.6 |
| Finger | 3896.6 ± 44.3 |

Table 1. EDR (Mean ± SD) to the radiation worker using the simplified geometry model at distance of 30 cm. from PMMA shielding apparatus and finger equivalent dose rate (contracted the knob)

The maximum EDR was at finger (3896.6 ± 44.3 µSvh-1), which touched the knob. According to 500 mSv ICRP finger dose limit of radiation worker, the calculated EDR was limited the SIRT for 128 patients per year assuming that the radiation worker performed 1 hour per procedure [22]. However, the knob contracted time was much lower in clinical routine. The contract times of 3, 5, 10 and 15 minutes were resulted the equivalent finger doses of 0.19, 0.32, 0.65 and 0.97 mSv per procedure, respectively. These were correlated to 2631, 1563, 769 and 515 cases per year, respectively. The EDR obtained from MC simulations can be reduced in practice by decreasing the contact time.

Study of Mrdja D et al evaluated the finger dose from 1- 3 GBq of 90Y bremsstrahlung radiation received by medical staff using Geant4 simulation [7].  The finger dose of 5 mSv (for 10 minutes per procedure) was stated which was estimated to 100 cases per year to reach the annual equivalent finger dose limit. The simulated finger equivalent dose from this study was higher than our results. Nevertheless, this further means that

the annual finger dose should be taken to account for SIRT procedure even the 90Y is well placed in the radiation shielding apparatus. Apart from finger EDR estimation, our simulation was calculated the body EDR of 18.88 ± 0.39 µSvh-1 for 3 GBq. Mrdja et al also reported the dose rate of 20 µSvh-1 for similar activity of 90Y placed within shielding. It is important to note that the inconsistent in finger dose might come from the dissimilarity finger model in simulation, shielding material and simulation code.

- Conclusions

This work demonstrated that 90Y bremsstrahlung radiation can be produced from the PMMA radiation shielding apparatus. Monte Carlo with MCNP5 can be used as a tool for evaluation of radiation safety aspect from bremsstrahlung photons in SIRT procedure. Prior to simulation, MCNP5 code was validated with 137Cs standard source and NaI(Tl) scintillation detector. The GEB was applied by using parameters from three gamma sources. The difference between simulation and experiment was well accepted.

For radiation safety aspect, the simulation result showed that the EDR to head, neck, body, and lower extremities to the radiation worker using simplified geometry model using data given in ICRP publication 89 was simulated with 90Y activity of 3 GBq, at distance of 30 cm.  The body EDR for radiation worker was much lower than the whole-body annual dose limit. However, the finger EDR can possibly reach the annual dose limit when radiation worker performed 2631, 1563, 769 and 515 cases per year with contract the knob 3, 5, 10 and 15 minutes per case respectively. Based on our findings, the finger EDR can be reduced in practice by decreasing the contact time. Future work with regard to additional protective equipment for SIRT procedure e.g., novel material for shielding apparatus should be examined. the finger EDR can be reduced in practice by decreasing the contact time. Future work with regard to additional protective equipment for SIRT procedure e.g., novel material for shielding apparatus should be examined.

- References

1. Townsend, A.R., et al., Selective internal radiation therapy for liver metastases from colorectal cancer. Cancer treatment reviews, (2016). 50: p. 148-154.

2. Jakobs, T.F., et al., Mid-term results in otherwise treatment refractory primary or secondary liver confined tumours treated with selective internal radiation therapy (SIRT) using 90 Yttrium resin-microspheres. European radiology, (2007). 17(5): p. 1320.

3. Lee, E.W., et al., Yttrium-90 selective internal radiation therapy with glass microspheres for hepatocellular carcinoma: current and updated literature review. Korean journal of radiology, (2016). 17(4): p. 472-488.

4. Italiano, A., L. Auditore, and E. Amato, Enhancement of radiation exposure risk from β-emitter radionuclides due to Internal Bremsstrahlung effect: A Monte Carlo study of 90Y case. Physica Medica, (2020). 76: p. 159-165.

5. Tijani, S., et al., Shielding of beta and bremsstrahlung radiation with transparent Bi2O3-B2O3-TeO2 glasses in therapeutic nuclear medicine. Journal of Radiological Protection, (2018). 38(3): p. N44.

6. Kim, Y.-C., et al., Radiation safety issues in Y-90 microsphere selective hepatic radioembolization therapy: possible radiation exposure from the patients. Nuclear medicine and molecular imaging, (2010). 44(4): p. 252-260.

7. Mrdja, D., et al., Monte Carlo simulation of beta particle-induced bremsstrahlung doses. J Radiol Prot, (2018). 38(1): p. 34-47.

8. Tuan, N.A. and C.V. Tao, Bremsstrahlung conversion efficiency and gamma-neutron generation from polypropylene, aluminum, iron, and lead bombarded by 10 MeV electrons. Nuclear Technology and Radiation Protection, (2020). 35(1): p. 36-41.

9. Team, M.C., MCNP A General Monte Carlo N-Particle Transport Code Version 5 Volume I: Overview and Theory. Los Alamos National Laboratory, Los Alamos, NM, LA-UR-03-1987, (2003).

10. Becquerel, L.N.H., Table de Radionucleides. 2014, Laboratoire National Henri Becquerel: Gif-sur-Yvette Cedex, France.

11. Hubbell, J.H., et al., Atomic form factors, incoherent scattering functions, and photon scattering cross sections. Journal of physical and chemical reference data, (1975). 4(3): p. 471-538.

12. Babaei, M., et al., Evaluation of Bremsstrahlung radiation dose in stereotactically radiocolloid therapy of cystic craniopharyngioma tumors with 32 P radio-colloid. Australasian physical & engineering sciences in medicine, (2018). 41(3): p. 697-711.

13. Eftekhari Zadeh, E., et al., Gaussian Energy Broadening Function of an HPGe Detector in the Range of 40 keV to 1.46 MeV. Journal of Experimental Physics, (2014). 2014: p. 623683.

14. Yttrium, S.-S., Microspheres [package insert]. Singapore Science Park SSM.

15. Valentin, J., Basic anatomical and physiological data for use in radiological protection: reference values: ICRP Publication 89. Annals of the ICRP, (2002). 32(3-4): p. 1-277.

16. Massey, J., Data for Protection against Ionizing Radiation from External Sources: Supplement to ICRP Publication 15. 1974, Taylor & Francis.

17. Salgado, C.M., et al., Validation of a NaI (Tl) detector's model developed with MCNP-X code. Progress in Nuclear Energy, (2012). 59: p. 19-25.

18. Haidar, H., F. Liu, and H. Yuan, Calculation of scintillation properties of Ø1″×1″ of the lanthanum bromide scintillation detector using MCNP simulation and experiment. Journal of Physics: Conference Series, (2018). 1053: p. 012063.

19. Fantínová, K., P. Fojtík, and I. Malátová, Monte Carlo simulation of the bremsstrahlung radiation for the measurement of an internal contamination with pure-beta emitters in vivo. Radiation Protection Dosimetry, (2016). 170(1-4): p. 354-358.

20. Al-Buriahi, M.S., et al., Mechanical and radiation shielding properties of tellurite glasses doped with ZnO and NiO. Ceramics International, (2020). 46(11, Part B): p. 19078-19083.

21. Kaur, P., K.J. Singh, and S. Thakur, Gamma-ray, Neutron and Beta Radiation Shielding Properties of Bi2O3-Li2O-MnO2-B2O3 Glasses. Journal of Physics: Conference Series, (2020). 1531: p. 012043.

22. Protection, R., ICRP publication 103. Ann ICRP, (2007). 37(2.4): p. 2.